## Test info
TI Stimulator information
Manufacturer
Model
Serial number
Testing information
Date
Operator
Test equipment

## Current output accuracy
Resistors H-Bridge |  | Channel 1 |  | Channel 2
Total resistance (kOhm) | 2
Sense resistor R1 (kOhm) | 1 | Setting |  | Setting
Sense resistor R2 (kOhm) | 1 | Frequency (Hz) |  | Frequency (Hz)
|  | Amplitude (mA) |  | Amplitude (mA)
Using Multimeter? | FALSE | * Set duration sufficently long for reading. |  | * Set duration sufficently long for reading.
|  | Measurement |  | Measurement
|  | Voltage on R1 (V) |  | Voltage on R2 (V)
|  | Calculated Current (mA) | 0 | Calculated Current (mA) | 0
|  | Current error (%) | #DIV/0! | Current error (%) | #DIV/0!
|  | Target amplitude error: < 5% |  | Target amplitude error: < 5%
|  | Frequency (Hz) |  | Frequency (Hz)
|  | Frequency error (Hz) | 0 | Frequency error (Hz) | 0
|  | Target frequency error: < 0.1 Hz |  | Target frequency error: < 0.1 Hz
Instructions
Perform the procedure on each channel individually, focusing on a single current at a time.
1. Connect the TI stimulator and measurement equipment to the resistor H-bridge (sense resistors 1 and 2).
2. Set & Apply the TI stimulation test parameters.
Default parameters:
· Frequency: 2000 Hz
· Amplitude: 2mA
· Duration: 10s (no ramping)
3. Read voltage on sense resistor & Log measurement voltage.
4. Calculate current from measured voltage.
5. Assess TI stimulator accuracy.
Additional tests (optional)
1. Current leakage: Set and apply 0 mA currents and validate that no current is applied.
2. Timing: Measure the duration and ramping periods and test against set values.

## Frequency response
Resistors H-Bridge |  | Channel 1 | [5 empty] | Channel 2
Total resistance (kOhm) | 2
Sense resistor R1 (kOhm) | 1 | Setting | [5 empty] | Setting
Sense resistor R2 (kOhm) | 1 | Set frequency (kHz) | 1 | 5 | 10 | 20 | 50 | Set frequency (kHz) | 1 | 5 | 10 | 20 | 50
|  | Amplitude (mA) | 2 | same for all frequencies |  |  |  | Amplitude (mA) | 2 | same for all frequencies
Using Multimeter? | FALSE | * Set duration sufficently long for reading. | [5 empty] | * Set duration sufficently long for reading.
|  | Measurement | [5 empty] | Measurement
|  | Voltage on R1 (V) | [5 empty] | Voltage on R1 (V)
|  | Calculated Current (mA) | 0 | 0 | 0 | 0 | 0 | Calculated Current (mA) | 0 | 0 | 0 | 0 | 0
|  | Current error (%) | -100 | -100 | -100 | -100 | -100 | Current error (%) | -100 | -100 | -100 | -100 | -100
|  | Target current error: < 5% | [5 empty] | Target current error: < 5%
Instructions
Perform the procedure on each channel individually, focusing on a single current at a time.
1. Connect the TI stimulator and measurement equipment to the resistor H-bridge (sense resistors 1 and 2).
2. Set & Apply the TI stimulation test parameters.
Default parameters:
· Amplitude: 2mA
· Duration: 10s (no ramping)
3. Read voltage on sense resistor & Log measurement voltage.
4. Calculate current from measured voltage.
5. Assess TI stimulator accuracy.

## Channel isolation
Resistors H-Bridge |  | Setting
Total resistance (kOhm) | 2 | Channel 1 |  | Channel 2
Sense resistor R1 (kOhm) | 1 | Frequency (Hz) |  | Frequency (Hz)
Sense resistor R2 (kOhm) | 1 | Amplitude (mA) |  | Amplitude (mA)
Sense resistor R3 (kOhm) | 0.068 | * Set duration sufficently long for reading.
Use Oscilloscope. |  | Measurement
|  | Channel 1 |  | Channel 2 |  | Crossing
|  | Max voltage on R1 (V) |  | Max Voltage on R2 (V) |  | Max Voltage on R3 (V)
|  | Min voltage on R1 (V) |  | Min Voltage on R2 (V) |  | Min Voltage on R3 (V)
|  | Calculated AM index (0-1) | #DIV/0! | Calculated AM index (0-1) | #DIV/0! | Calculated AM index (0-1) | #DIV/0!
|  | Target AM index < 0.05 |  | Target AM index < 0.05 |  | Target AM index > 0.95
Instructions
1. Connect the TI stimulator and measurement equipment to the resistor H-bridge (sense resistors 1 ,2, and 3).
2. Set & Apply the TI stimulation test parameters.
Default parameters:
· Frequencies: 2000 Hz, 2010Hz
· Amplitudes: 2mA, 2mA
· Duration: 10s (no ramping)
3. Read voltage on both sense resistors & Log measurement voltages.
4. Calculate amplitude modulation index for each channel.
5. Assess TI stimulator accuracy.